## Slide 1: Development of Spherical Neutron Polarimetry Device
Tianhao Wang
Advisor: Prof. Roger Pynn

## Slide 2: Outline
Introduction to Spherical Neutron Polarimetry technology (SNP)
Current existing SNP devices (CryoPAD, MuPAD)
Research goal
Previous research experiment result
Next step research plan
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 3: General Description of SNP
  Polarized neutron diffraction experiment method
  Makes maximum usage of neutron spin manipulation
  Introduce and detect neutron spin polarization at any direction
* M. Janoshek (2007)
X: Along Q
Z: Vertical
Y: Form a Cartesian coordinate
Analyzer
& Detector
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 4: Basic formulae
Basic formula for polarized neutron scattering developed by Blume (1963) and Maleeve (1999)
  : Total scattering cross section : Final polarization
  : Initial polarization
  : Magnetic scattering vector component perpendicular to (x direction)
Nuclear
Incoherent
Magnetic
Interference
Chiral
Nuclear & incoherent
Magnetic
Interference
Chiral

## Slide 5: From formula to experiment - Example
MnWO4 - Pure coherent, magnetic scattering
Want to measure the magnetic and chiral scattering term

## Slide 6: Existing SNP device – CryoPAD at ILL since 20 years ago
Developed since over 20 years ago.
pioneer work, massive improvement have been made
nowadays most successful SNP device
Can be mounted on Triple axis spectrometer
- Not limited on scattering angle
- Not limited on elastic scattering
Independent sample control
- temperature, external electric field etc.
Expensive, Labor consuming
    - cryogenic liquid drive

## Slide 7: Schematics of CryoPAD
  Nutator: shift neutron spin from parallel to propagation to perpendicular
  Precession coil: shift neutron spin via Larmor precesion
B
B

## Slide 8: Alternative design - MuPAD
Technique highlights
Cryogenic free
  - less cost, less work
Can be mounted on Triple axis spectrometer
- available for all kinds of scattering experiment
Polarization controlled by two precession field
Need very careful engineering on Precession coil windings and resident field removal
    - High requirement on engineering support, also expensive

## Slide 9: MuPAD Schematics
MuPAD (prototype build in 2004) at FRM-II in German
Design and developed by M. Janoschek

## Slide 10: Research Goal
Spherical Neutron Polarimetry is a very powerful experimental method to detect material structure and properties.
Currently no similar device exists in U.S.
Research target:
Develop a device that could fulfill spherical neutron polarimetry experiment method, based on available neutron sources in U.S. (LENS at IU, HFIR at ORNL, NIST? Los Alamos?)
Features of the device:
Use High Tc superconducting material (YBCO) for magnetic field shielding.
Use closed cycle cryogenic stats
Not bound to one single source
Taking advantage of Time of Flight source ?
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 11: Research Plan
Build and test precession field shielded by YBCO (Spin nutator)
realize and calibrate Spherical neutron spin manipulation
Small angle neutron scattering based SNP experiment
Build Fully functional SNP device
Condensed matter experiment
(Further Improvement)
  (Time of Flight method, advanced sample manipulation……)
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 12: First Step of research
Neutron spin nutator
  Crucial component of the whole device
  Test of High Tc YBCO superconducting material as B field shield
Guide field on rotation stage
(Spin rotator)
Front precession field
(spin nutator)
Back precession field
(spin nutator)
Zero field chamber
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 13: Nutator component and Assemble Procedure
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 14: Spin Nutator Simulation using Radia
(Excellent Work from colleague Fankang Li)
Field Integral mapping
Wavelength: 4 Ang
Current: 4 A
180 deg rotation
Dephasing caused by
Inhomogeneity:
(20mm beam size)
< 0.1%
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 15: Cryogenic system setup
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 16: Real Picture and Difficulties
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 17: Beam Setup
Guide Field
Cryogenic system
Analyzer
Detector
Front Slit
Front Pi Flipper
Vacuum Chamber
Front & Back Electro Magnets
Nutator chamber
Back Slit
Notes: Pictures clockwise from the top – Front end view showing permanent magnet guide field and cryo-cup location, vies of instrument, detail of up stream pi flipper (not used except for setup) and detail of guide field/electromagnets around cup.

## Slide 18: Experiment Setup
Simple measurement:
Time of flight source – precession angle vs. neutron wavelength
Shifting precession coil current - Precession angle vs. current
Incoming Guide Field
Outgoing Guide Field
Precession Field
Analyzer Field
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 19: Polarization vs. Wavelength plot with different precession coil currents
Wavelength
4 Angstrom
Polarization
8 Angstrom

## Slide 20: Plot of P/PI=0 and fitting to function
Wavelength
4 Angstrom
8 Angstrom

## Slide 21: Fitting Data and the Beam size Comparison
Wavelength
Depolarization: <2% with 20mm beam size
<0.4% with 15mm beam size
<0.1% with 10mm beamsize

## Slide 22: Choosing a Specific Wavelength and plot polarization through different current
Compared to CryoPAD precession: 123.65deg/(A [Ang]), CryoCUP has less precession power: 11.52deg /(A [Ang]).
Current

## Slide 23: Experiment summary:
Through the experiment test, it is feasible to use YBCO superconducting film as a magnetic field shield to isolate precession region from outer field
By adjusting the current going through coil, we are able to control the precession angle.
The dephasing factor of <2% is within acceptable range so that the nutator is suitable for next step development
The thermal anchor of cryogenic system is not good, which limited the highest current to 4A
The homogeneity of field generated by Helmholtz coil is not ideal, also the field strength generated is not strong enough
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 24: Next step development
Research Steps:
Neutron spin nutator
Spherical neutron spin manipulation
Spherical neutron polarimetry with SANS
Fully functional SNP device
Condensed matter experiment
Further development
  (Time of Flight method, advanced sample manipulation……)
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 25: Next step Device development
Guide field on rotation stage
(Spin rotator)
Front precession field
(spin nutator)
Back precession field
(spin nutator)
Zero field chamber
Analyzing direction
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 26: Capacity of Spin nutator for SANS experiment
For a typical 90 nutation, 3.438 deg difference is generated, which causes
~ 0.17% depolarization
0.3deg – 5deg
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 27: Next Step Device – What experiment?
MnSi – Candidant for SANS experiment
Curie temperature: 29K
Below Tc: single chirality spiral
structure
Propagation vector:
Small Angel scattering experiment
Done by both CryoPAD and MuPAD
Phase Diagram
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 28: Research Plan
Research Steps:
Neutron spin nutator
Spherical neutron spin manipulation
Spherical neutron polarimetry with SANS
Fully functional SNP device
Condensed Matter Experiment
(Further Improvement – Time of Flight method)
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 29: Next step development
Fully functional SNP device on a Optical Platform?
Spin rotator
Spin nutator
Spin rotator
Spin nutator
Analyzer
Detector
Detachable
Mu-Metal Shield
Sample
Optical Platform
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 30: Summary
Spherical Neutron Polarimetry is a very powerful experimental method to detect material.
Develop a device that could fulfill spherical neutron polarimetry experiment method, based on available neutron sources in U.S.
The first step of research, which is making spin nutator has finished, the nutator made is suitable for further development
Still a lot of work to do.
Notes: [1] Spherical neutron polarimetry with MuPAD - M. Janoschek (2007)

## Slide 31
Slit
Slit
V Coil
V Coil
Wollaston Prism
Source
Polarizaer &
Monochromater
Analyzer &
Detector